What should our project do?

The goal of this project is to build a system to measure the number of heartbeats per minute of any person who want to measure the speed of his blood flow that fits comfortably into daily living.

After putting the finger in the specific place, the number of the heartbeats will be shown on the LCD automatically.

Work principle of our project:

When the person put his finger inside the hollow, indeed his finger become between a high intensity LED 4200mcd and a responsive resistor (photo resistor ), On every heartbeat reaches his finger it makes light isolation on the resistor, this mean that the heartbeat will prevent the rays from the intensity LED to reach the responsive resistor, and this makes the responsive resistor  change its resistance to be lower than initial value when the high intensity light is on it directly, so this brings bigger current to the resistor branch, this current will enter the Operational Amplifier OpAmp(LM358) that in his tern amplify the current to max scale of 5 volts which will be translated to pulses, each beat represents one pulse, then these sequence of pulses enters to the microcontroller On pin A0 That capture the pulse happening and also the next pulse and take time between them, in the system 3 pulses is enough to give pulse indication speed but for the accuracy we tacked 5 pulses and then we computed the  time for the one pulse to make sure that’s no wrong reading in the pulses that arrives and show the number of heartbeats per minute on the LCD that connected to the pic16f877A and the Multiplexed Display that connected to the pic16f84A.

The figure below shows the Signals detected and processed by our project:

Structure of our project:

Regulator Circuit:

- As known this part of project is to transform 8-15VDC 100mA input

To 5VDC output to provide safety of the hardware pieces.

Heartbeat Sensor Circuit:

- This circuit aims to produce suitable heartbeat pulse to enter the

microcontrollers.

- As seen in the pictures above, there is a hollow to put a finger inside it, inside the hollow there is a high intensity LED and responsive resistor to detect each heartbeat.

- The produced signal is sent to the LM 358 (OpAmp) to make the inputs suitable to be detected on the microcontroller.

- There is two variable resistors to control with trigger level. ((the current that enters the OpAmp to be amplified)) and sensitivity ((to what extend resistor catch changes in light)).

- There is also two LEDs connected to the OpAmp output indicates the pulse sequence that enters the controller and to show the accuracy of the produced signal.

Pic16F84 Microcontroller Circuit:

- This microcontroller control the reached signals and do its calculations to show the result output which is the number of heartbeats per minute on the multiplexed display.

Multiplexed Display Circuit:

- This multiplexed display is to show the number of heartbeats per minute.

- It is controlled by the pic16F84 microcontroller.

- It consists of three common cathode Seven Segment LEDs.

Pic16F877 Microcontroller and LCD Circuit:

- This microcontroller control the reached signals by calculate the time between the pulses that comes over Pin A0 and do its calculations to show the result output which is the number of heartbeats per minute on the LCD.

Complete Circuit:

- A number of pictures taken to the project from many angles is shown below:

Programming in our project:

- As mentioned above this project contains two microcontrollers.

- Pic16F84A microcontroller is programmed by 'assembly'.

- Pic16F877A microcontroller is programmed by 'C' language.

- This is the program of pic16F84 microcontroller.

list p=16f84,r=hex	;microcontroller & base

include	"reg16c84.inc"	;register memory mapping file

include "project.inc"	;heartbeat definitions

;****************************************************************

;global variables

;****************************************************************

count		equ	0x0c		;general purpose counter

display3	equ	0x0d		;data for 7-segment display 100's

display2	equ	0x0e		;data for 7-segemnt display 10's

display1	equ	0x0f		;data for 7-segment display 1's

beatlsb		equ	0x10		;heartbeat counter,lsb

beatmsb		equ	0x11		;heartbeat counter,msb

aargb0		equ	0x12		;dividend,lsb

aargb1		equ	0x13		;dividend,msb

bargb0		equ	0x14		;divisor,lsb

bargb1		equ	0x15		;divisor,msb

remb0		equ	0x16		;remainder,lsb

remb1		equ	0x17		;remainder,msb

timera		equ	0x18		;general purpose timer

timerb		equ	0x19		;general purpose timer

timerc		equ	0x1a		;general purpose timer

timerd		equ	0x1b		;general purpose timer

;****************************************************************

;reset vector

;****************************************************************

reset:	org	0x00			;reset vector address

goto	start			;start program execution

start:	org 	0x06			;start of program

goto	Initialisation

;****************************************************************

; LED DISPLAY DECODER FROM BINARY TO 7-SEGMENT

;****************************************************************

LedDecoder:

addwf	pcl,f								;W   DISPLAY

retlw	LED_SEGa|LED_SEGb|LED_SEGc|LED_SEGd|LED_SEGe|LED_SEGf		;0 - "0"

retlw	LED_SEGb|LED_SEGc						;1 - "1"

retlw	LED_SEGa|LED_SEGb|LED_SEGd|LED_SEGe|LED_SEGg			;2 - "2"

retlw	LED_SEGa|LED_SEGb|LED_SEGc|LED_SEGd|LED_SEGg			;3 - "3"

retlw	LED_SEGb|LED_SEGc|LED_SEGf|LED_SEGg				;4 - "4"

retlw	LED_SEGa|LED_SEGc|LED_SEGd|LED_SEGf|LED_SEGg			;5 - "5"

retlw	LED_SEGa|LED_SEGc|LED_SEGd|LED_SEGe|LED_SEGf|LED_SEGg		;6 - "6"

retlw	LED_SEGa|LED_SEGb|LED_SEGc					;7 - "7"

retlw	LED_SEGa|LED_SEGb|LED_SEGc|LED_SEGd|LED_SEGe|LED_SEGf|LED_SEGg	;8 - "8"

retlw	LED_SEGa|LED_SEGb|LED_SEGc|LED_SEGd|LED_SEGf|LED_SEGg		;9 - "9"

retlw	LED_BLANK							;10- " "

;****************************************************************

;Name		: Initialisation

;Description	: system initialisation and mainloop

;Inputs		: none

;Outputs	: none

;****************************************************************

Initialisation:

;system initialisation

;port A and port B initialisation

bcf	status,rp1		;select bank1 register

bsf	status,rp0

movlw	PORTA_CONFIG1		;porta configuration

movwf	trisa

movlw	PORTB_CONFIG1		;portb configuration

movwf	trisb

bcf	status,rp0		;select bank0 registers

movlw	0x00			;disable all displays

movwf	porta

;****************************************************************

;Name		: Mainloop

;Description	: measure and display heartbeat

;Inputs		: none

;Outputs	: none

;****************************************************************

Mainloop:

call	Heartbeat		;measure heartbeat

call	Convert			;convert to bpm

call	Display			;show on display

goto	Mainloop		;endless loop

;****************************************************************

;Name		: Heartbeat

;Description	: measures the time between two heartbeat signals.

;		: Triggered on rising-edge of signal, then measures

;		: amount of milli-seconds before next rising-edge.

;Inputs		: none

;Outputs	: beatlsb,beatmsb = length between two heartbeats

;		: in milli-seconds

;****************************************************************

Heartbeat:

btfss	porta,HEARTBEAT_SENSOR	;wait until signal goes high

goto	Heartbeat

call	Delay150		;wait 150msec - wait until signal has gone low

clrf	beatmsb			;preset beat registers with 150msec

movlw	0x96

movwf	beatlsb

HL1:	call	Delay1			;wait 1msec

incf	beatlsb,w		;increment beatlsb by 1msec

movwf	beatlsb

btfsc	status,z		;lsb rollover?

incf	beatmsb,f		;yes -increment msb

btfss	porta,HEARTBEAT_SENSOR	;keep looping till signal goes back high

goto	HL1

return

;****************************************************************

;Name		: Convert

;Description	: convert the measured heartbeat from milliseconds

;		: to beats-per-minute. (maybe there's a quicker way!)

;Inputs		: beatlsb, beatmsb

;Outputs	: display1,display2,display3 = contain hearbeat

;		: in bcd format

;****************************************************************

Convert:

;calculate bpm = 60000/beat (lsb and msb)

;Make dividend 60000 == 0xEA60

movlw	0xea		;msb

movwf	aargb1

movlw	0x60		;lsb

movwf	aargb0

;Get measured heartbeat values, as divisor

movf	beatmsb,w	;msb

movwf	bargb1

movf	beatlsb,w	;lsb

movwf	bargb0

clrf	remb0		;initialise remainder registers

clrf	remb1

; do an unsigned 16-bit by 16-bit division

movlw	0x10

movwf	count

Lu16:	rlf	aargb1,w

rlf	remb0,f

rlf	remb1,f

movf	bargb0,w

subwf	remb0,f

movf	bargb1,w

btfss	status,c

incfsz	bargb1,w

subwf	remb1,f

btfsc	status,c

goto	uok16

movf	bargb0,w

addwf	remb0,f

movf	bargb1,w

btfsc	status,c

incfsz	bargb1,w

addwf	remb1,f

bcf	status,c

uok16:	rlf	aargb0,f

rlf	aargb1,f

decfsz	count,f

goto	Lu16

;convert the lsb result of the previous division to unpacked bcd

movlw	0x08			;an 8-bit division

movwf	count

movlw	0x0a			;divide the previous lsb result by 10

movwf	bargb0

clrf	remb0			;initialise remainder

Lu8:	rlf	aargb0,w

rlf	remb0,f

movf	bargb0,w

subwf	remb0,f

btfsc	status,c

goto 	uok8

addwf	remb0,f

bcf	status,c

uok8:	rlf	aargb0,f

decfsz	count,f

goto	Lu8

movlw	0x0a			;3rd bcd digit - 100's is blank(if not changed further down code!)

movwf	display3

movf	remb0,w

movwf	display1		;1st bcd digit - 1's

movf	aargb0,w

movwf	display2		;2nd bcd digit - 10's(if not changed further down code!)

sublw	0x09			;is the quotient >=10?

bnc	$+2			;yes -quotient is >=10, so divide by 10 again!

return				;go back

movlw	0x08			;an 8-bit division

movwf	count

movlw	0x0a			;divide the previous lsb result by 10

movwf	bargb0

clrf	remb0			;initialise remainder

Lu81:	rlf	aargb0,w

rlf	remb0,f

movf	bargb0,w

subwf	remb0,f

btfsc	status,c

goto 	uok81

addwf	remb0,f

bcf	status,c

uok81:	rlf	aargb0,f

decfsz	count,f

goto	Lu81

movf	remb0,w

movwf	display2		;2nd bcd digit - 10's

movf	aargb0,w

movwf	display3		;3rd bcd digit - 100's

return				;go back

;****************************************************************

;Name		: Display

;Description	: updates display with current measurement

;Inputs		: display1 - 7-segment display 1's

;		: display2 - 7-segment display 10's

;		: display3 - 7-segment display 100's

;Outputs	: none

;****************************************************************

Display:

movf	display1,w		;convert bcd to 7-segment

call	LedDecoder

movwf	display1

movf	display2,w

call	LedDecoder

movwf	display2

movf	display3,w

call	LedDecoder

movwf	display3

movlw	0x4			;update display for a little while

movwf	timerc

DL2:	movlw	0xff

movwf	timerd

DL1:	clrf	porta

bsf	porta,LED_DIGIT1

movf	display1,w

movwf	portb

call	Delay1

clrf	porta

bsf	porta,LED_DIGIT2

movf	display2,w

movwf	portb

call	Delay1

clrf	porta

bsf	porta,LED_DIGIT3

movf	display3,w

movwf	portb

call	Delay1

clrf	porta

decfsz	timerd,f

goto	DL1

decfsz	timerc,f

goto	DL2

return				;go back

;****************************************************************

;Name		: Delay

;Description	: causes a delay of 1msec

;Inputs		: none

;Outputs	: none

;****************************************************************

Delay1:

movlw	0x02			;number of wait states

movwf	timerb

L2:

movlw	0xa5			;number of wait states

movwf	timera

L3:	decfsz	timera,f

goto	L3

decfsz	timerb,f

goto	L2

return				;go back

;****************************************************************

;Name		: Delay

;Description	: causes a delay of 150msec

;Inputs		: none

;Outputs	: none

;****************************************************************

Delay150:

movlw	0xc3			;number of wait states

movwf	timerb

L4:

movlw	0xff			;number of wait states

movwf	timera

L5:	decfsz	timera,f

goto	L5

decfsz	timerb,f

goto	L4

return				;go back

end

- This is the C program of pic16F877 microcontroller.

#include "pulse.h"

#include<LCD.C>

void print_pulse();

int pulse_n;

int i,j;

//int pulse_nnn;

double result;

void noob();

void main()

{

setup_adc_ports(NO_ANALOGS);

setup_adc(ADC_OFF);

setup_psp(PSP_DISABLED);

setup_spi(FALSE);

setup_counters(RTCC_INTERNAL,RTCC_DIV_1);

setup_timer_1(T1_DISABLED);

setup_timer_2(T2_DISABLED,0,1);

setup_comparator(NC_NC_NC_NC);

setup_vref(FALSE);

lcd_init();

pulse_n=0;

result=0;

restart_wdt();

i=0;

while(1)

{

result=0;  // time for one pulse

noob();   // to test if there no finger putted

restart_wdt();

for(j=0;j<5;j++) //for take 5 pulse value and compute the time for pulse

{

pulse_n=0;

delay_ms(250);

restart_wdt();

pulse_n=250;

restart_wdt();

while(!input(PIN_A0))

{

restart_wdt();

delay_ms(1);

pulse_n++;

restart_wdt();

}

result=result+pulse_n;   // result is the time for 5 pulses

}

result=result/5;    // result is the average time for for the pulse

result=60000/result;   // pulse / min

restart_wdt();

if(result>60 && result<100)   // regect the error pulses

{

restart_wdt();

print_pulse();   // print function

}

}

}

void print_pulse()

{

restart_wdt();

printf(LCD_PUTC, "\f");

restart_wdt();

printf(LCD_PUTC, "pulse=");

restart_wdt();

printf(LCD_PUTC, "%ld",result);

delay_ms(300);

restart_wdt();

//result=0;

// pulse_n=0;

return;

}

void noob() // test if no fingure putted on the sensor room

{

while(!input(PIN_A0))

{

i++;

delay_ms(1);

if(i>3000)  // sensor room empty for 3 seconds

{

restart_wdt();

printf(LCD_PUTC, "\f");

printf(LCD_PUTC, "PUT FINGER");

restart_wdt();

i=0;

}

}

restart_wdt();

return;

}

# Text boxes

AN-Najah National University

Computer Engineering Department

Hardware Project

Project:

Heartbeats Monitor.

Made by:

Mohammad Zorba.

Mohammad Zahlan.